## HDR Quarterly QA
HDR Brachytherapy
Quarterly QA Worksheet
|  | Date Completed | Completed By
1 | Exit dose calibration
2 | Source position accuracy - part 1
3 | Source position accuracy - part 2 autoradiograph
4 | Source calibration
Decay chart complete & posted
5 | Daily QA tests
Battery check in Flashlight
6 | Catheter obstruction test
7 | Power failure recovery
8 | Source data verification
9 | Mechanical integrity of applicators & transfer tubes
10 | Records
Radiation survey forms complete
Daily pre-treatment QA records complete

## Source Receipt Survey
192Ir HDR SOURCE RECEIPT RADIATION SURVEY FORM
Date of Receipt: | 2018-06-06 |  |  | CCSEO Batch Number: |  |  | #40
192Ir Source Information from Manufacturer
Serial Number: | D85E2661 |  |  | Calibration Date: |  |  | 2018-05-24
Source Strength: | 426.3 GBq |  |  | & Time: |  |  | 00:00:00 CET
RADIATION SURVEY UPON NEW SOURCE RECEIPT
(NOTE: 1 mSv = 100 mR and 1 mR = 10 mSv/h)
MEASURING SURVEY METER
Make/Model: | Fluke 451P |  | S/N: | 1543
TRANSPORT BUCKET/DOCUMENTATION | [6 empty] | ACCEPTABLE?
Tranport Bucket is NOT damaged
Essential source RECEIPT documets present (source handling procedure, Certificate for Sealed Sources, Source Data Plate, Delivery note)
Source RETURN Shipment documets present
|  |  |  | LOCATION |  | DOSE RATE (mSv/h) | ACCEPTABLE?
MAXIMUM EXPOSURE RATES AROUND THE TRASPORT BUCKET (mSv/h):
Max dose rate at surface of transport bucket (DS) | [5 empty] | 80
Max dose rate at 1 meter from surface of transport bucket (D1m) | [5 empty] | 5
Corridor on north side of bunker | [5 empty] | Background
Contamination Measurements (on a wipe on an area of 300 cm²) | [5 empty] | Background
NOTE: Acceptable values - Ds ≤ 500 mSv/h, D1m ≤ 10 mSv/h, Contamination ≤ 0.4 Bq/cm²)
STORAGE AND INVENTORY
Source bucket storage location:
Source information entered in the 192Ir Source Inventroy Control Form?
Signature of Surveyor: |  | Dr. T. Olding |  |  | Date of Survey: | 2018-06-07
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
HDR Bunker Closet
Radioisotope Lab
Co-60 Bunker
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
HDR Bunker Closet
Radioisotope Lab
Co-60 Bunker
Yes
No

## Exit Dose Calibration
Department of Medical Physics, Cancer Centre of South Eastern Ontario @ KGH
EXIT Src Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: | 39 |  | TDU: | Flexitron B (FT 00107)
Date and Time: | 2018-06-07 12:02 |  | Flexisource No.: | NLF 01 D85E2431
HDR - 1000 Well-Type Ionization Chamber Serial Number: |  |  | A021025 (245cc) | Date of Calibration: March 1, 2014 | 2018-05-15
Calibration Factor (CWC): | 462100 | Gy m2h-1A-1 (or U/A) = |  | 462.1 | mGy m2h-1/nA-1 (or U/A)
Electrometer Model: Max 4000 Serial Number: E021531
Calibration Factor (CE): |  | 1 | nA/Rdg
Bias Voltage: 302 V
Temperature - Pressure Correction Factor: | [5 empty] | Determination of Sweet spot
Temperature (°C): | 22.8 |  | 758.2 | CTP : | 1.0051
|  |  | Final Measurements (at the location of sweetspot) |  |  | full voltage | [6 empty] | 1/2 voltage
Leakage (nA) |  | Rdgs (nA) | Source Position (mm) | Rdgs (nA) |  | 248 | 249 | 250 | 251 | 252 | 253 | 255 | 253
Before Calibration |  |  | 249 | -38.952 |  | -38.944 | -38.952 | -38.953 | -38.944 | -38.931
0 |  |  |  | -38.95 |  | -38.945 | -38.95 | -38.95 | -38.949 | -38.93
After Calibration |  |  |  | -38.948 |  | -38.944 | -38.948 | -38.946 | -38.937 | -39.929
0 |  |  |  | -38.953 |  | -38.945 | -38.953 |  | -38.938
Average leakage (nA) |  |  |  | -38.948 |  | -38.946 | -38.948 |  | -38.935
0 |  |  |  | -38.954 |  | -38.945 | -38.954
|  |  | Average | -38.9508
Pion Measurements: (Charge 30 sec each) |  |  |  |  | Average
Reading (full voltage) Q1: -300 | -2699.7 | -2699.7 | -2699.6 | nC | -2699.6667
Reading (half voltage) Q2: -150 | -2690.3 | -2693.2 | -2693.3 | nC | -2692.2667
Pion = [(4/3) - (Q1/3Q2)]-1 = |  | 1.0009
Source Strength Calculations: Sk (mGy m2h-1) = R x CTP x CE x Cwc x Pion | [5 empty] | -38.9448 | -38.9508 | -38.9497 | -38.9406 | -39.2633
R = | Avg. Reading at full voltage at position of max. response |  |  | -38.9508
CTP = | Temperature and pressure correction factor |  |  | 1.0051
CE = | Electrometer Calibration Factor |  |  | 1
Cwc = | Well-chambered calibration factor |  |  | 462.1
Pion = | Recombination Corr. Factor |  |  | 1.0009
Calculated SK = | 18107.4034 | (mGy m2h-1) | Date: | 2018-06-07 12:02
= | 164.6128 | GBq = | 4.449 | Curies | at Calibration
Manufaturer's Calibration:
Activity stated by manufacturer | 447.5 |  | Manufacturer's source strength on
Date and Time (CET): | 2018-02-21 20:58
Date and Time(Local): | 2018-02-21 14:58
Decay Factor | 0.37017 |  | 18221.4763 | U (mGy m2h-1)
Activity on Console |  |  | 18228 | U (mGy m2h-1)
Activity at present, Manufacturer's calibration |  |  | 18221.4763 | U (mGy m2h-1)
Activity Calculated at present, CCSEO calibration |  |  | 18107.4034 | U (mGy m2h-1)
% Variation WRT console (%VAR1)* |  |  | -0.66601 | % | Acceptable
% Variation WRT Manufacturer's calibration (%VAR2)* |  |  | -0.62998 | % | Acceptable
Signature: | T.Olding |  |  | Date: | 2018-06-07
Note on A2: Enter The Readings at YELLOW Colors
Note on B4: cjoshi: enter date and time in correct format e.g 27 Aug 2008 17:00
Note on B35: cjoshi: cjoshi: enter date and time in correct format e.g 15 Aug 2008 11:28
Note on B36: cjoshi: cjoshi: enter date and time in correct format e.g 15 Aug 2008 11:28
Note on B37: cjoshi: time difference between Holland and Canada accounted for

## New Source Calibration
Department of Medical Physics, Cancer Centre of South Eastern Ontario @ KGH
NEW Src Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: | 40 |  | TDU: | Flexitron B (FT 00107)
Date and Time: | 2018-06-08 14:00 |  | Flexisource No.: | NLF 01 D85E2661
HDR - 1000 Well-Type Ionization Chamber Serial Number: |  |  | A021025 (245cc) | Date of Calibration: March 1, 2014 | 2018-05-15
Calibration Factor (CWC): | 462100 | Gy m2h-1A-1 (or U/A) = |  | 462.1 | mGy m2h-1/nA-1 (or U/A)
Electrometer Model: Max 4000 Serial Number: E021531
Calibration Factor (CE): |  | 1 | nA/Rdg
Bias Voltage: 302 V
Temperature - Pressure Correction Factor: | [5 empty] | Determination of Sweet spot
Temperature (°C): | 23.1 |  | 760.8 | CTP : | 1.0027
|  |  | Final Measurements (at the location of sweetspot) |  |  | full voltage | [6 empty] | 1/2 voltage
Leakage (nA) |  | Rdgs (nA) | Source Position (mm) | Rdgs (nA) |  | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 250
Before Calibration |  |  | 250 | -86.711 |  | -86.662 | -86.707 | -86.711 | -86.701 | -86.691 | -86.68 | -86.629
0 |  |  |  | -86.709 |  | -86.663 | -86.705 | -86.709 | -86.701 | -86.692 | -86.674 | -86.629
After Calibration |  |  |  | -86.71 |  | -86.664 | -86.71 | -86.71 | -86.704 | -86.692 | -86.674 | -86.626
0 |  |  |  | -86.712 |  | -86.663 | -86.707 | -86.712 | -86.705 | -86.689 | -86.677 | -86.626
Average leakage (nA) |  |  |  | -86.712 |  | -86.665 | -86.703 | -86.712 | -86.709 | -86.69
0 |  |  |  | -86.714 |  | -86.666 | -86.705 | -86.714 | -86.709 | -86.688
|  |  | Average | -86.7113
Pion Measurements: (Charge 30 sec each) |  |  |  |  | Average
Reading (full voltage) Q1: -300 | -2699.7 | -2699.7 | -2699.6 | nC | -2699.6667
Reading (half voltage) Q2: -150 | -2690.3 | -2693.2 | -2693.3 | nC | -2692.2667
Pion = [(4/3) - (Q1/3Q2)]-1 = |  | 1.0009
Source Strength Calculations: Sk (mGy m2h-1) = R x CTP x CE x Cwc x Pion | [5 empty] | -86.6638 | -86.7062 | -86.7113 | -86.7048 | -86.6903 | -86.6763 | -86.6275
R = | Avg. Reading at full voltage at position of max. response |  |  | -86.7113
CTP = | Temperature and pressure correction factor |  |  | 1.0027
CE = | Electrometer Calibration Factor |  |  | 1
Cwc = | Well-chambered calibration factor |  |  | 462.1
Pion = | Recombination Corr. Factor |  |  | 1.0009
Calculated SK = | 40213.1948 | (mGy m2h-1) | Date: | 2018-06-08 14:00
= | 365.5745 | GBq = | 9.8804 | Curies | at Calibration
Manufaturer's Calibration:
Activity stated by manufacturer | 426.3 |  | Manufacturer's source strength on
Date and Time (CET): | 2018-05-24
Date and Time(Local): | 2018-05-23 19:00
Decay Factor | 0.86224 |  | 40432.8701 | U (mGy m2h-1)
Activity on Console |  |  | 40437 | U (mGy m2h-1)
Activity at present, Manufacturer's calibration |  |  | 40432.8701 | U (mGy m2h-1)
Activity Calculated at present, CCSEO calibration |  |  | 40213.1948 | U (mGy m2h-1)
% Variation WRT console (%VAR1)* |  |  | -0.55655 | % | Acceptable
% Variation WRT Manufacturer's calibration (%VAR2)* |  |  | -0.54628 | % | Acceptable
Signature: | TO/JDJ |  |  | Date: | 2018-06-08
Note on A2: Enter The Readings at YELLOW Colors
Note on B4: cjoshi: enter date and time in correct format e.g 27 Aug 2008 17:00
Note on B35: cjoshi: cjoshi: enter date and time in correct format e.g 15 Aug 2008 11:28
Note on B36: cjoshi: cjoshi: enter date and time in correct format e.g 15 Aug 2008 11:28
Note on B37: cjoshi: time difference between Holland and Canada accounted for

## New Source Survey
NEW 192Ir SOURCE RADIATION SURVEY FORM
Date of Receipt: | 2018-06-08 |  |  | CCSEO Batch Number: |  |  | 40
192Ir Source Information from Manufacturer
Serial Number: | D85E2661 |  |  | Calibration Date: |  |  | 2018-05-24
Source Strength: | 426.3 GBq |  |  | & Time: |  |  | 00:00:00 CET
RADIATION SURVEYS FOR THE FLEXITRON UNIT WITH THE NEW SOURCE
(NOTE: 1 mSv = 100 mR and 1 mR = 10 mSv/h)
MEASURING SURVEY METER
Make/Model: | Fluke 451 |  | S/N: | 1543
|  |  |  | LOCATION |  | DOSE RATE (mSv/h) | ACCEPTABLE?
MAXIMUM EXPOSURE RATES NEAR THE FLEXITRON UNIT WITH SOURCE FULLY RETRACTED (mSv/h):
At the surface of the unit (allowed max ≤ 500 mSv/h or 50 mR/h) | [5 empty] | 2.7
At 1 meter from the surface of the unit (allowed max ≤ 10 mSv/h or 1 mR/h) | [5 empty] | 0.2
MAXIMUM EXPOSURE RATES AT ADJACENT AREAS WITH SOURCE EXTENDED (mSv/h):
HDR Suite Entrance door (@ surface only) | [5 empty] | 0.07
HDR Control Console | [5 empty] | 0.05
Storage Rooms (southwest) | [5 empty] | 0.06
Entrance Corridor (north) | [5 empty] | 0.13
0.3
1.1
SOURCE LEAKAGE TEST RESULTS ON SOURCE CERTIFICATE CHECKED
Source is leak free and the surface contamination activity is ~200 Bq (~5 nCi)
Signature of Surveyor: |  | TO/JDJ |  |  | Date of Survey: | 2018-06-08
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No

## Decay Chart
Iridium-192 HDR Brachytherapy Source (Flexisource) Decay Chart
CCSEO Batch Number: |  | 40 |  |  |  |  | Manufacturer's Calibration: |  |  | 426.3 | GBq |  | Flexisource No: | NLF 01 D85E2661
Date of source installation @ CCSEO: |  |  | 2018-06-08 14:00 |  |  |  | Date & Time (CET) of Calibration |  |  | 2018-05-24
Day | Date | Day | *Source Strength | Day | Date | Day | *Source Strength | Day | Date | Day | *Source Strength | Day | Date | Day | *Source Strength
Number | (At 12:00) |  | (U) | Number | (At 12:00) |  | (U) | Number | (At 12:00) |  | (U) | Number | (At 12:00) |  | (U)
1 | 2018-06-08 12:00 | 2018-06-08 12:00 | 40464.5088 | 31 | 2018-07-08 12:00 | 2018-07-08 12:00 | 30533.8599 | 61 | 2018-08-07 12:00 | 2018-08-07 12:00 | 23040.3539 | 91 | 2018-09-06 12:00 | 2018-09-06 12:00 | 17385.8761
2 | 2018-06-09 12:00 | 2018-06-09 12:00 | 40086.4737 | 32 | 2018-07-09 12:00 | 2018-07-09 12:00 | 30248.6008 | 62 | 2018-08-08 12:00 | 2018-08-08 12:00 | 22825.102 | 92 | 2018-09-07 12:00 | 2018-09-07 12:00 | 17223.4505
3 | 2018-06-10 12:00 | 2018-06-10 12:00 | 39711.9703 | 33 | 2018-07-10 12:00 | 2018-07-10 12:00 | 29966.0066 | 63 | 2018-08-09 12:00 | 2018-08-09 12:00 | 22611.861 | 93 | 2018-09-08 12:00 | 2018-09-08 12:00 | 17062.5424
4 | 2018-06-11 12:00 | 2018-06-11 12:00 | 39340.9657 | 34 | 2018-07-11 12:00 | 2018-07-11 12:00 | 29686.0526 | 64 | 2018-08-10 12:00 | 2018-08-10 12:00 | 22400.6123 | 94 | 2018-09-09 12:00 | 2018-09-09 12:00 | 16903.1375
5 | 2018-06-12 12:00 | 2018-06-12 12:00 | 38973.4272 | 35 | 2018-07-12 12:00 | 2018-07-12 12:00 | 29408.714 | 65 | 2018-08-11 12:00 | 2018-08-11 12:00 | 22191.3371 | 95 | 2018-09-10 12:00 | 2018-09-10 12:00 | 16745.2219
6 | 2018-06-13 12:00 | 2018-06-13 12:00 | 38609.3223 | 36 | 2018-07-13 12:00 | 2018-07-13 12:00 | 29133.9664 | 66 | 2018-08-12 12:00 | 2018-08-12 12:00 | 21984.017 | 96 | 2018-09-11 12:00 | 2018-09-11 12:00 | 16588.7815
7 | 2018-06-14 12:00 | 2018-06-14 12:00 | 38248.6191 | 37 | 2018-07-14 12:00 | 2018-07-14 12:00 | 28861.7856 | 67 | 2018-08-13 12:00 | 2018-08-13 12:00 | 21778.6338 | 97 | 2018-09-12 12:00 | 2018-09-12 12:00 | 16433.8027
8 | 2018-06-15 12:00 | 2018-06-15 12:00 | 37891.2856 | 38 | 2018-07-15 12:00 | 2018-07-15 12:00 | 28592.1476 | 68 | 2018-08-14 12:00 | 2018-08-14 12:00 | 21575.1694 | 98 | 2018-09-13 12:00 | 2018-09-13 12:00 | 16280.2717
9 | 2018-06-16 12:00 | 2018-06-16 12:00 | 37537.2906 | 39 | 2018-07-16 12:00 | 2018-07-16 12:00 | 28325.0287 | 69 | 2018-08-15 12:00 | 2018-08-15 12:00 | 21373.6058 | 99 | 2018-09-14 12:00 | 2018-09-14 12:00 | 16128.1751
10 | 2018-06-17 12:00 | 2018-06-17 12:00 | 37186.6027 | 40 | 2018-07-17 12:00 | 2018-07-17 12:00 | 28060.4053 | 70 | 2018-08-16 12:00 | 2018-08-16 12:00 | 21173.9253 | 100 | 2018-09-15 12:00 | 2018-09-15 12:00 | 15977.4995
11 | 2018-06-18 12:00 | 2018-06-18 12:00 | 36839.191 | 41 | 2018-07-18 12:00 | 2018-07-18 12:00 | 27798.2541 | 71 | 2018-08-17 12:00 | 2018-08-17 12:00 | 20976.1103 | 101 | 2018-09-16 12:00 | 2018-09-16 12:00 | 15828.2315
12 | 2018-06-19 12:00 | 2018-06-19 12:00 | 36495.025 | 42 | 2018-07-19 12:00 | 2018-07-19 12:00 | 27538.5521 | 72 | 2018-08-18 12:00 | 2018-08-18 12:00 | 20780.1433 | 102 | 2018-09-17 12:00 | 2018-09-17 12:00 | 15680.358
13 | 2018-06-20 12:00 | 2018-06-20 12:00 | 36154.0744 | 43 | 2018-07-20 12:00 | 2018-07-20 12:00 | 27281.2762 | 73 | 2018-08-19 12:00 | 2018-08-19 12:00 | 20586.0072 | 103 | 2018-09-18 12:00 | 2018-09-18 12:00 | 15533.866
14 | 2018-06-21 12:00 | 2018-06-21 12:00 | 35816.309 | 44 | 2018-07-21 12:00 | 2018-07-21 12:00 | 27026.404 | 74 | 2018-08-20 12:00 | 2018-08-20 12:00 | 20393.6847 | 104 | 2018-09-19 12:00 | 2018-09-19 12:00 | 15388.7426
15 | 2018-06-22 12:00 | 2018-06-22 12:00 | 35481.6992 | 45 | 2018-07-22 12:00 | 2018-07-22 12:00 | 26773.9129 | 75 | 2018-08-21 12:00 | 2018-08-21 12:00 | 20203.159 | 105 | 2018-09-20 12:00 | 2018-09-20 12:00 | 15244.9751
16 | 2018-06-23 12:00 | 2018-06-23 12:00 | 35150.2154 | 46 | 2018-07-23 12:00 | 2018-07-23 12:00 | 26523.7806 | 76 | 2018-08-22 12:00 | 2018-08-22 12:00 | 20014.4133 | 106 | 2018-09-21 12:00 | 2018-09-21 12:00 | 15102.5506
17 | 2018-06-24 12:00 | 2018-06-24 12:00 | 34821.8285 | 47 | 2018-07-24 12:00 | 2018-07-24 12:00 | 26275.9852 | 77 | 2018-08-23 12:00 | 2018-08-23 12:00 | 19827.4309 | 107 | 2018-09-22 12:00 | 2018-09-22 12:00 | 14961.4567
18 | 2018-06-25 12:00 | 2018-06-25 12:00 | 34496.5095 | 48 | 2018-07-25 12:00 | 2018-07-25 12:00 | 26030.5047 | 78 | 2018-08-24 12:00 | 2018-08-24 12:00 | 19642.1953 | 108 | 2018-09-23 12:00 | 2018-09-23 12:00 | 14821.681
19 | 2018-06-26 12:00 | 2018-06-26 12:00 | 34174.2297 | 49 | 2018-07-26 12:00 | 2018-07-26 12:00 | 25787.3177 | 79 | 2018-08-25 12:00 | 2018-08-25 12:00 | 19458.6903 | 109 | 2018-09-24 12:00 | 2018-09-24 12:00 | 14683.2111
20 | 2018-06-27 12:00 | 2018-06-27 12:00 | 33854.9608 | 50 | 2018-07-27 12:00 | 2018-07-27 12:00 | 25546.4026 | 80 | 2018-08-26 12:00 | 2018-08-26 12:00 | 19276.8997 | 110 | 2018-09-25 12:00 | 2018-09-25 12:00 | 14546.0349
21 | 2018-06-28 12:00 | 2018-06-28 12:00 | 33538.6747 | 51 | 2018-07-28 12:00 | 2018-07-28 12:00 | 25307.7382 | 81 | 2018-08-27 12:00 | 2018-08-27 12:00 | 19096.8074 | 111 | 2018-09-26 12:00 | 2018-09-26 12:00 | 14410.1402
22 | 2018-06-29 12:00 | 2018-06-29 12:00 | 33225.3434 | 52 | 2018-07-29 12:00 | 2018-07-29 12:00 | 25071.3035 | 82 | 2018-08-28 12:00 | 2018-08-28 12:00 | 18918.3977 | 112 | 2018-09-27 12:00 | 2018-09-27 12:00 | 14275.5151
23 | 2018-06-30 12:00 | 2018-06-30 12:00 | 32914.9393 | 53 | 2018-07-30 12:00 | 2018-07-30 12:00 | 24837.0777 | 83 | 2018-08-29 12:00 | 2018-08-29 12:00 | 18741.6547 | 113 | 2018-09-28 12:00 | 2018-09-28 12:00 | 14142.1477
24 | 2018-07-01 12:00 | 2018-07-01 12:00 | 32607.4352 | 54 | 2018-07-31 12:00 | 2018-07-31 12:00 | 24605.0401 | 84 | 2018-08-30 12:00 | 2018-08-30 12:00 | 18566.5629 | 114 | 2018-09-29 12:00 | 2018-09-29 12:00 | 14010.0262
25 | 2018-07-02 12:00 | 2018-07-02 12:00 | 32302.8039 | 55 | 2018-08-01 12:00 | 2018-08-01 12:00 | 24375.1702 | 85 | 2018-08-31 12:00 | 2018-08-31 12:00 | 18393.1069 | 115 | 2018-09-30 12:00 | 2018-09-30 12:00 | 13879.1392
26 | 2018-07-03 12:00 | 2018-07-03 12:00 | 32001.0186 | 56 | 2018-08-02 12:00 | 2018-08-02 12:00 | 24147.448 | 86 | 2018-09-01 12:00 | 2018-09-01 12:00 | 18221.2713 | 116 | 2018-10-01 12:00 | 2018-10-01 12:00 | 13749.4749
27 | 2018-07-04 12:00 | 2018-07-04 12:00 | 31702.0527 | 57 | 2018-08-03 12:00 | 2018-08-03 12:00 | 23921.8532 | 87 | 2018-09-02 12:00 | 2018-09-02 12:00 | 18051.0412 | 117 | 2018-10-02 12:00 | 2018-10-02 12:00 | 13621.0219
28 | 2018-07-05 12:00 | 2018-07-05 12:00 | 31405.8799 | 58 | 2018-08-04 12:00 | 2018-08-04 12:00 | 23698.3659 | 88 | 2018-09-03 12:00 | 2018-09-03 12:00 | 17882.4013 | 118 | 2018-10-03 12:00 | 2018-10-03 12:00 | 13493.7691
29 | 2018-07-06 12:00 | 2018-07-06 12:00 | 31112.474 | 59 | 2018-08-05 12:00 | 2018-08-05 12:00 | 23476.9666 | 89 | 2018-09-04 12:00 | 2018-09-04 12:00 | 17715.337 | 119 | 2018-10-04 12:00 | 2018-10-04 12:00 | 13367.7051
30 | 2018-07-07 12:00 | 2018-07-07 12:00 | 30821.8092 | 60 | 2018-08-06 12:00 | 2018-08-06 12:00 | 23257.6357 | 90 | 2018-09-05 12:00 | 2018-09-05 12:00 | 17549.8335 | 120 | 2018-10-05 12:00 | 2018-10-05 12:00 | 13242.8188
* NOTE: Source Air Kerma Strength (U) must be within ±100 U of the source strength value on the treatment delivery unit (TDU)
|  |  |  | 1 U = 1cGy.cm2.h-1 = 1mGy.m2.h-1 |  |  |  | 1 Curie = 4070 U |  | 1 GBq = 110 U
Formulas:
Day = IF([Date]="","",[Date])
Day = [Date]
Date = IF([Date, row above]="","",[Date, row above]+1)
Note on C10: Halsall, Tracy: insert date and time in correct format eg. 01 Dec 2015 12:00 (table will automatically populate)

## Daily Safety & QA
HDR Brachytherapy
Pretreatment Safety & QA Tests
Date: |  | June 8 2018
DAILY SAFETY & QA TESTS
1 | Startup | P
2 | Emergency equipment | P
3 | Survey meter | P
4 | Visual communication | P
5 | Date, time & source strength in treatment unit
| Expected Activity (U) | 40444
| Current Activity (U) | 40444
| Calculated dwell time of 60s | 60.3798
| Calculated dwell time of 15s | 15.0949
| Dwell times accurate | P
6 | Emergency off
| on TDU | P
| on wall inside treatment room | P
| on TCP (no radiation) | P
| on TCP (with radiation) | P
7 | Warning lights | P
8 | Room radiation monitors | P
9 | Treatment interrupt | P
10 | Door interlock / last person out | P
11 | Key switch | P
12 | Transit tube misconnect interlock | P
13 | Source positional accuracy
| Channel 1 - 200mm | 200
| Channel 3 - 180mm | 180
| Channel 3 - 220mm | 220
| Channel 5 - 200mm | 200
14 | Dwell time accuracy
| Expected dwell time (s) | 60.3798
| Measured dwell time (s) | 61.06
SIGNATURE |  | JD Jutras & T Olding

## Power Failure Recovery
HDR Brachytherapy
Power Failure Recovery
"Terminate" QA Plan | Plan Name | Terminate_08062018
| Channel | 1
| Dwell position | 200
| Dwell time (sec) | 30
Treatment Details (Actualized) | Radiation Time (sec) | 30.2
| TRAK | 0.339
| AK Strength | 40421
TCP at time of interruption | Last Channel | 1
| Last Dwell | 200
| Remaining Time (sec) | 15.4
TCP at time of reset | Remaining Time (sec) | 15.4 | PASS
Treatment Details (Delivered) for "Terminated" Plan | Radiation time (sec) | 14.8 | PASS
| TRAK | 0.166
"Recovery" QA Plan | Plan Name | Recovery_08062018
| Channel | 1
| Dwell position | 200
| Calculated dwell time for 40700U | 15.2944
Treatment Details (Actualized) | Radiation time (sec) | 15.5 | PASS | entered 15.4s into recovery plan instead of 15.3s as specified
| TRAK | 0.174 | PASS
Treatment Details (Delivered) for "Recovery" Plan | Radiation time (sec) | 15.5 | PASS
| TRAK | 0.174 | PASS
Performed by: | TO/JDJ

## Applicators & Transfer Tubes
HDR Brachytherapy
Mechanical Integrity of Applicators & Transfer Tubes
|  |  | Date Completed | Completed By
1 | Flexible implant transfer tubes | (Physics)
2 | Lung Catheter | (Physics)
3 | Cervix-CT/MR Ring Applicator
| Set 1 | (RT)
| Set 2 | (RT)
4 | Vaginal Vault Stainless Steel Applicator
| Set 1 | (RT)
| Set 2 | (RT)
5 | Bonvoisin-Gerard Esophageal Applicator
| 6mm | (RT)
| 8mm | (RT)
| 10mm | (RT)

## Source Return Survey
192Ir HDR SOURCE RETURN RADIATION SURVEY FORM
Date of Return: | 2018-06-18 |  |  | CCSEO Batch Number: |  |  | 39
192Ir Source Information from Manufacturer
Serial Number: | NLF 01 D85E2431 |  |  | Calibration Date: |  |  | 2018-02-21
Source Strength: | 447GBq |  |  | & Time: |  |  | 20.58 CET
RADIATION SURVEY UPON SOURCE RETURN
(NOTE: 1 mSv = 100 mR and 1 mR = 10 mSv/h)
MEASURING SURVEY METER
Make/Model: | Fluke 451 |  | S/N: | 1543
TRANSPORT BUCKET/DOCUMENTATION | [6 empty] | ACCEPTABLE?
Tranport Bucket is NOT damaged
Correct source activity in GBq is written on the two "II-Yellow" labels (that are marked "Return" and pasted on opposite sides of the transport bucket)
Correct source activity in GBq is written on FlexiSource Return Form
3 Copies of the Shippers Declaration Form for Dangerous Goods is attached
|  |  |  | LOCATION |  | DOSE RATE (mSv/h) | ACCEPTABLE?
MAXIMUM EXPOSURE RATES AROUND THE TRASPORT BUCKET (mSv/h):
Max dose rate at surface of transport bucket (DS) | [5 empty] | 46
Max dose rate at 1 meter from surface of transport bucket (D1m) | [5 empty] | 1.7
Corridor on north side of bunker1 | [5 empty] | n/a
Other areas
Contamination Measurements (on a wipe on an area of 300 cm²) | [5 empty] | backgnd.
NOTE: Acceptable values - Ds ≤ 500 mSv/h, D1m ≤ 10 mSv/h, Contamination ≤ 0.4 Bq/cm²)
Transport Index2 (TI) = [D1m (in mSv/h) ¸ 10] = | [5 empty] | 0.2
NOTES:
1 Measurements outside the Transport Bucket storage area shoub be ≤ 2.5 mSv/h
2 TI must be rounded up to the 1st decimal place (e.g. 0.11 must be rounded up to 0.2). TI should be < 0.6; DS should not be > 200 mSv/h and contamination activity should be ≤ 0.4 Bq/cm²
Signature of Surveyor: |  | Christopher B. Peters |  |  | Date of Survey: | 2018-06-14
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No
Yes
No