## 18th December 2008
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 2
Date and Time: | 2008-12-17 12:00 |  | Flexisource No.: NLF 01 D85A-091
Manufaturer's Calibration:
Activity stated by manufacturer | 358.46 | GBq | Manufacturer's source strenght on
Date and Time: | 2008-12-08 13:16 |  | 2008-12-17 12:00
Decay Factor | 0.91729 |  | 36169.4293 | U (mGy m2h-1)
Activity on Console |  |  | 35883.11 | U (mGy m2h-1)
Activity at present, Manufacturer's calibration |  |  | 36169.4293 | U (mGy m2h-1)
Note on A2: Enter The Readings at YELLOW Colors
Note on B4: cjoshi: enter date and time in correct format e.g 27 Aug 2008 17:00
Note on B7: cjoshi: cjoshi: enter date and time in correct format e.g 15 Aug 2008 11:28
Note on B8: cjoshi: time difference between Holland and Canada accounted for

## 4th Feb 2010
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 6
Date and Time: | 2010-02-04 12:00 |  | Flexisource No.: NLF 01 D85A-235
Manufaturer's Calibration:
Activity stated by manufacturer | 428.34 | GBq | Manufacturer's source strenght on
Date and Time: | 2010-01-21 15:51 |  | 2010-02-04 12:00
Decay Factor | 0.87612 | 47117.4 | 41280.6261 | U (mGy m2h-1)
Activity on Console |  |  | 35883.11 | U (mGy m2h-1)
Activity at present, Manufacturer's calibration |  |  | 41280.6261 | U (mGy m2h-1)
Note on B8: cjoshi: time difference between Holland and Canada accounted for

## September 10 2010
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 8
Date and Time: | 2010-08-28 14:21 | (Holland) | Flexisource No.: NLF 01 D85A-339
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2010-08-28 08:21 | (Here)
Activity stated by manufacturer | 442.81 | GBq | 48709.1 | U (mGy m2h-1)
Local Calibration: | Date and Time: | 2010-09-09 15:06
| CTP | 1.0045
| Pion | 1.0008
| Calibration Factor | 467.6319 | U/A (mGy m2h-1nA-1)
| Average Current | -92.192 | nA
Measured Activity |  |  | 43340.5685 | U (mGy m2h-1)
| Decay Factor | 0.89112
Expected activity from manufacturer |  |  | 43405.6823 | U (mGy m2h-1)
|  |  | -0.15024 | % difference
Activity set on Console | Date and Time: | 2010-09-09 15:48
Console Activity |  |  | 43400.31 | U (mGy m2h-1)
| Decay Factor | 0.89088
Expected activity from manufacturer |  |  | 43393.8008 | U (mGy m2h-1)
|  |  | 0.014998 | % difference

## December 6 2010
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 9
Date and Time: | 2010-11-16 13:48 | (Holland) | Flexisource No.: NLF 01 D85A-388
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2010-11-16 07:48 | (Here)
Activity stated by manufacturer | 412.8 | GBq | 45408 | U (mGy m2h-1) |  | ((T+273) x 760)/(Px295) =
Local Calibration: | Date and Time: | 2010-12-06 18:08
| CTP | 1.0215
| Pion | 1.0008
| Calibration Factor | 467.6319 | U/A (mGy m2h-1nA-1)
| Average Current | -78.255 | nA
Measured Activity |  |  | 37411.2219 | U (mGy m2h-1)
| Decay Factor | 0.8255
Expected activity from manufacturer |  |  | 37484.2541 | U (mGy m2h-1)
|  |  | -0.19521 | % difference
Activity set on Console | Date and Time: | 2010-11-16 07:48
Console Activity |  |  | 45408 | U (mGy m2h-1)
| Decay Factor | 1
Expected activity from manufacturer |  |  | 45408 | U (mGy m2h-1)
|  |  | 0 | % difference

## Feb 24 2011
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 10
Date and Time: | 2011-02-12 11:52 | (Holland) | Flexisource No.: NLF 01 D85A-587
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2011-02-12 05:52 | (Here)
Activity stated by manufacturer | 447.38 | GBq | 49211.8 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2011-02-24 13:51 |  |  |  | Tempurature | 21.3
| CTP | 1.0016 |  |  |  | Pressure | 757
| Pion | 1.0009 |  |  |  | CTP | 1.0016
| Calibration Factor | 467.6319 | U/A (mGy m2h-1nA-1)
| Average Current | -93.494 | nA
Measured Activity |  |  | 43829.2984 | U (mGy m2h-1)
| Decay Factor | 0.89069
Expected activity from manufacturer |  |  | 43832.5008 | U (mGy m2h-1)
|  |  | -0.0073066 | % difference
Activity set on Console | Date and Time: | 2011-02-24 13:48
Console Activity |  |  | 43831.07 | U (mGy m2h-1)
| Decay Factor | 0.89071
Expected activity from manufacturer |  |  | 43833.3579 | U (mGy m2h-1)
|  |  | -0.0052199 | % difference

## Sept 7 2011
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 10
Date and Time: | 2011-08-19 08:10 | (Holland) | Flexisource No.: NLF 01 D85A-587
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2011-08-19 02:10 | (Here)
Activity stated by manufacturer | 422.96 | GBq | 46525.6 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2011-09-07 16:35 |  |  |  | Tempurature | 20.7
| CTP | 0.99559 |  |  |  | Pressure | 760
| Pion | 1.0008 |  |  |  | CTP | 0.99559
| Calibration Factor | 467.6319 | U/A (mGy m2h-1nA-1)
| Average Current | -82.751 | nA
Measured Activity |  |  | 38557.2994 | U (mGy m2h-1)
| Decay Factor | 0.83195
Expected activity from manufacturer |  |  | 38707.1632 | U (mGy m2h-1)
|  |  | -0.38868 | % difference
Activity on Console | Date and Time: | 2011-09-07 16:35
Console Activity |  |  | 38727.01 | U (mGy m2h-1)
| Decay Factor | 0.83195
Expected activity from manufacturer |  |  | 38707.1632 | U (mGy m2h-1)
|  |  | 0.051248 | % difference

## Dec 2 2011
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 13
Date and Time: | 2011-11-09 07:31 | (Holland) | Flexisource No.: NLF 01 D85A-587
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2011-11-09 01:31 | (Here)
Activity stated by manufacturer | 457.66 | GBq | 50342.6 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2011-12-01 18:00 |  |  |  | Tempurature | 20.3
| CTP | 0.99033 |  |  |  | Pressure | 763
| Pion | 1.0008 |  |  |  | CTP | 0.99033
| Calibration Factor | 467.6319 | U/A (mGy m2h-1nA-1)
| Average Current | -87.466 | nA
Measured Activity |  |  | 40538.6977 | U (mGy m2h-1)
| Decay Factor | 0.8082
Expected activity from manufacturer |  |  | 40686.9112 | U (mGy m2h-1)
|  |  | -0.36561 | % difference
Activity on Console | Date and Time: | 2011-12-01 18:00
Console Activity |  |  | 40674.63 | U (mGy m2h-1)
| Decay Factor | 0.8082
Expected activity from manufacturer |  |  | 40686.9112 | U (mGy m2h-1)
|  |  | -0.030194 | % difference

## Sept 10 2012
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 16
Date and Time: | 2012-08-20 21:33 | (Holland) | Flexisource No.: NLF 01 D85A-587
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2012-08-20 15:33 | (Here)
Activity stated by manufacturer | 393.26 | GBq | 43258.6 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2012-09-07 16:30 |  |  |  | Tempurature | 21.2
| CTP | 1.0072 |  |  |  | Pressure | 752.5
| Pion | 1.0007 |  |  |  | CTP | 1.0072
| Calibration Factor | 469 | U/A (mGy m2h-1nA-1)
| Average Current | -77.19 | nA
Measured Activity |  |  | 36489.2993 | U (mGy m2h-1)
| Decay Factor | 0.84423
Expected activity from manufacturer |  |  | 36520.4032 | U (mGy m2h-1)
|  |  | -0.085241 | % difference
Activity on Console | Date and Time: | 2012-09-10 08:14
Console Activity |  |  | 35615.38 | U (mGy m2h-1)
| Decay Factor | 0.82345
Expected activity from manufacturer |  |  | 35621.3527 | U (mGy m2h-1)
|  |  | -0.01677 | % difference

## June 18 2013
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 20
Date and Time: | 2013-05-28 19:52 | (Holland) | Flexisource No.: NLF 01 D85B-161
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2013-05-28 13:52 | (Here)
Activity stated by manufacturer | 408.76 | GBq | 44963.6 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2013-06-18 14:30 |  |  |  | Tempurature | 22
| CTP | 1.0074 |  |  |  | Pressure | 754.4
| Pion | 1.0008 |  |  |  | CTP | 1.0074
| Calibration Factor | 469 | U/A (mGy m2h-1nA-1)
| Average Current | -78.676 | nA
Measured Activity |  |  | 37202.6883 | U (mGy m2h-1)
| Decay Factor | 0.8209
Expected activity from manufacturer |  |  | 36910.3964 | U (mGy m2h-1)
|  |  | 0.78567 | % difference
Activity on Console | Date and Time: | 2013-06-18 14:30
Console Activity |  |  | 36903.8 | U (mGy m2h-1)
| Decay Factor | 0.8209
Expected activity from manufacturer |  |  | 36910.3964 | U (mGy m2h-1)
|  |  | -0.017874 | % difference

## Dec 16 2013
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 22
Date and Time: | 2013-11-26 20:28 | (Holland) | Flexisource No.: NLF 01 D85B-161
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2013-11-26 14:28 | (Here)
Activity stated by manufacturer | 395 | GBq | 43450 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2013-12-16 14:20 |  |  |  | Tempurature | 22.4
| CTP | 1.004 |  |  |  | Pressure | 758
| Pion | 1.0008 |  |  |  | CTP | 1.004
| Calibration Factor | 469 | U/A (mGy m2h-1nA-1)
| Average Current | -77.279 | nA
Measured Activity |  |  | 36417.8661 | U (mGy m2h-1)
| Decay Factor | 0.82888
Expected activity from manufacturer |  |  | 36015.0525 | U (mGy m2h-1)
|  |  | 1.1061 | % difference
Activity on Console | Date and Time: | 2013-12-16 14:20
Console Activity |  |  | 36011.64 | U (mGy m2h-1)
| Decay Factor | 0.82888
Expected activity from manufacturer |  |  | 36015.0525 | U (mGy m2h-1)
|  |  | -0.009476 | % difference

## Sept 5 2014
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 25
Date and Time: | 2014-08-20 11:45 | (Holland) | Flexisource No.: NLF 01 D85E0220
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2014-08-20 05:45 | (Here)
Activity stated by manufacturer | 465.61 | GBq | 51217.1 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2014-09-04 16:30 |  |  |  | Tempurature | 22.4
| CTP | 1.0055 |  |  |  | Pressure | 756.9
| Pion | 1.0008 |  |  |  | CTP | 1.0055
| Calibration Factor | 463.1 | U/A (mGy m2h-1nA-1)
| Average Current | -94.519 | nA
Measured Activity |  |  | 44045.8259 | U (mGy m2h-1)
| Decay Factor | 0.86502
Expected activity from manufacturer |  |  | 44303.9865 | U (mGy m2h-1)
|  |  | -0.58612 | % difference
Activity on Console | Date and Time: | 2014-09-04 16:30
Console Activity |  |  | 44298.12 | U (mGy m2h-1)
| Decay Factor | 0.86502
Expected activity from manufacturer |  |  | 44303.9865 | U (mGy m2h-1)
|  |  | -0.013243 | % difference

## Sept 7 2017
Department of Medical Physics, Cancer Centre of Southe Eastern Ontario @ KGH
Calibration of 192Ir HDR Source (Flexisource)
CCSEO Batch No.: 27
Date and Time: | 2017-08-23 04:03 | (Holland) | Flexisource No.: NLF 01 D85E0220
Manufacturer's Calibration: | Manufacturer's source strength on |  |  | 2017-08-22 22:03 | (Here)
Activity stated by manufacturer | 455.5 | GBq | 50105 | U (mGy m2h-1)
((T+273) x 760)/(Px295)
Local Calibration: | Date and Time: | 2017-09-07 16:30 |  |  |  | Tempurature | 22
| CTP | 1.011 |  |  |  | Pressure | 751.7
| Pion | 1.0006 |  |  |  | CTP | 1.011
| Calibration Factor | 463.1 | U/A (mGy m2h-1nA-1)
| Average Current | -91.261 | nA
Measured Activity |  |  | 42755.2593 | U (mGy m2h-1)
| Decay Factor | 0.86242
Expected activity from manufacturer |  |  | 43211.6687 | U (mGy m2h-1)
|  |  | -1.0675 | % difference
Activity on Console | Date and Time: | 2017-09-07 16:30
Console Activity |  |  | 43230 | U (mGy m2h-1)
| Decay Factor | 0.86242
Expected activity from manufacturer |  |  | 43211.6687 | U (mGy m2h-1)
|  |  | 0.042404 | % difference